## Fig 3-1 - Nomenclature
FIG. 3-1
Nomenclature
C | = | Pitot tube flow coefficient | Fpb | = | pressure base factor applied to change the base pressure to 14.73 psia
C' | = | the product of multiplying all orifice correction factors | Fpm | = | pressure base factor applied to meter volumes to correct to standard pressure. See API Manual of Petroleum Measurement Standards, Chapter 12, Section 2.
Cpl | = | liquid pressure correction factor. Correction for the change in volume resulting from application of pressure. Proportional to the liquid compressibility factor, which depends upon both relative density and temperature. | Fpv | = | supercompressibility factor required to correct for deviation from the ideal gas laws = √(1/Z)
Cps | = | correction factor for effect of pressure on steel. See API Manual of Petroleum Measurement Standards, Chapter 12, Section 2. | Fs | = | steam factor
Cg | = | gravity correction factor for orifice well tester to change from a gas relative density of 0.6 | Fsl | = | orifice slope factor
Ctl | = | liquid temperature correction factor. Proportional to the thermal coefficient which varies with density and temperature | Ftb | = | temperature base factor. To change the temperature base from 60 ºF to another desired base
Cts | = | correction factor for effect of temperature on steel | Ftf | = | Flowing temperature factor to change from the assumed flowing temperature of 60 °F to the actual flowing temperature
Cu | = | velocity of sound in the gas non-flowing condition | Ftm | = | temperature correction factor applied to displacement meter volumes to correct to stadard temperature. See API Manual of Petroleum Measurement Standards, Chapter 12, Section 2
d | = | orifice diameter, in. | G, G1 | = | relative density (specific gravity) at 60 °F
D | = | internal pipe diameter of orifice meter run or prover section, in. | Gf | = | relative density (specific gravity) at flowing temperature
DL | = | Minimum downstream meter tube length, in. | H | = | pressure, inches of mercury
Dp | = | the difference between the flowing pressure and the equilibrium vapor pressure of the liquid. | hm | = | differential pressure measured across the orifice plate in inches of mercury at 60 ºF
Du | = | Diameter of the meter bore. | hw | = | differential pressure measured across the orifice plate in inches of water at 60 ºF
ρ | = | flowing fluid density, lb/ft3. | √(hwPf) | = | pressure extension. The square root of the differential pressure times the square root of the absolute static pressure
e | = | orifice edge thickness, in. | IM | = | indicated mass, mass measured at flowing conditions without correction for meter performance, counts / K. Also the difference between opening and closing meter readings.
E | = | orifice plate thickness, in. | IV | = | indicated volume, volume measured at flowing conditions without correction for meter performance, counts / K. Also the difference between opening and closing meter readings.
Em | = | modulus of elasticity for steel [(30)(106)] psi | k | = | ratio of the specific heat at constant pressure to the specific heat at constant volume
F | = | liquid compressibility factor | K | = | a numerical constant. Pulses generated per unit volume or mass through a turbine, positive displacement, coriolis or ultrasonic meter
Fa | = | orifice thermal expansion factor. Corrects for the metallic expansion or contraction of the orifice plate. Generally ignored between 0° and 120 °F | SEP | = | Separation distance between elbows, in. (See Fig. 3-9)
Fc |  | orifice calculation factor | SPU | = | Signal Processing Unit, the electronic microprocessor system of the multi-path ultrasonic meter.
Fg | = | relative density (specific gravity) factor applied to change from a specific gravity of 1.0 (air) to the specific gravity of the flowing gas | t1 | = | transit time from upstream transducer to downstream transducer.
Fgt | = | gravity-temperature factor for liquids | t2 | = | transit time from downstream transducer to upstream transducer.
Fn | = | numeric conversion factor | Tb | = | base temperature, °F
Fna | = | units conversion factor for pitot tubes | Tf | = | flowing temperature, °F
Key | = | Fn (Fc + Fsl) = orifice factor | Tmeas | = | reference temperature of the orifice plate bore, °F
L | = | distance between upstream and downstream transducer. | UL | = | Minimum upstream meter tube length, in. (See Fig. 3-9)
LTB | = | Length of tube bundle, in flow conditioner, in. (See fig. 3-3) | UL1 | = | See Fig. 3-9
MF | = | meter factor, a number obtained by dividing the actual quantity of liquid passed through the meter during proving by the quantity indicated or registered by the meter. | UL2 | = | See Fig. 3-9
N | = | number of whole meter pulses per single proving run | UM | = | ultrasonic meter for measuring gas flow rates
Navg | = | average number of pulses or interpolated pulses for proving runs that satisfy the repeatability requirements | V | = | velocity of flowing gas
P | = | pressure, psia. | VOS | = | velocity of sound
Pb | = | base pressure, psia | X | = | axial distance separating transducers
Pf1, Pf2 | = | static pressure at either the upstream(1) or downtream(2) pressure tap, psia | Y | = | expansion factor to compensate for the change in density as the fluid passes through an orifice.
Pp | = | operating or observed pressure, psig. | YCR | = | critical flow constant
Ps | = | pressure at which the base volume of a meter prover was determined, usually 0 psig. | Z | = | compressibility factor
P1/P2 | = | pressure ratio | β | = | ratio of the orifice or throat diameter to the internal diameter of the meter run, dimensionless
Q | = | gas flow rate, cu ft/day | ρ | = | density, lb/cu ft or lb/gal
Qh | = | rate of flow, std. cu ft/hr or gal./hr
Rh | = | maximum differential range, in. of water
Rp | = | maximum pressure range of pressure spring, psi
S | = | square of super compressibility

## Example 3-1
Example 3-1 -- Size an orifice plate in gas service | [6 empty] | Application 3-1 -- Size an orifice plate in gas service
Given Data: | [6 empty] | Given Data:
Line Size, D | = | 4.026 | in |  |  |  | Line Size, D | = | 4.026 | in
Flange Taps | [6 empty] | Flange Taps
Relative Density | = | 0.7 |  |  |  |  | Relative Density | = | 0.7
Flowing Temperature | = | 100 | °F |  |  |  | Flowing Temperature | = | 100 | °F
Flowing Pressure | = | 75 | psia |  |  |  | Flowing Pressure | = | 75 | psia
Flow Rate | = | 14200 | cu ft/hr |  |  |  | Flow Rate | = | 14200 | cu ft/hr
|  | 14.73 | psia, 60 ºF | [5 empty] | 14.73 | psia, 60 ºF
Desired Differential | = | 50 | in. of water |  |  |  | Desired Differential | = | 50 | in. of water
To determine the approximate orifice size required, the corresponding Keyg (natural gas) is calculated using appropriate terms of Ep. 3-4; | [6 empty] | To determine the approximate orifice size required, the corresponding Keyg (natural gas) is calculated using appropriate terms of Ep. 3-4;
Qh | = | Keyg • Ftf • Fg • (hw • Pf)1/2 |  |  | from Equation 3-4 |  | Qh | = | Keyg • Ftf • Fg • (hw • Pf)1/2 |  |  | from Equation 3-4
Rearranging, | [6 empty] | Rearranging,
Keyg | = | Qh / (Ftf • Fg • (hw • Pf)1/2) |  |  |  | Eq 3-13 | Keyg | = | Qh / (Ftf • Fg • (hw • Pf)1/2) |  |  |  | Eq 3-13
Ftf | = | [520 / (460 + Tf)]1/2 |  | = | 0.96362 | Fig 3-4 | Ftf | = | [520 / (460 + Tf)]1/2 |  | = | 0.96362 | Fig 3-4
Fg | = | [1.000/G]1/2 |  | = | 1.1952 | Fig 3-4 | Fg | = | [1.000/G]1/2 |  | = | 1.1952 | Fig 3-4
Keyg | = | 14200 |  | = | 201.3326 |  | Keyg | = | 14200 |  | = | 201.3326
|  | 0.964 (1.195) (50 • 75)1/2 | [6 empty] | 0.964 (1.195) (50 • 75)1/2
Select 1" orifice from Fig 3-16, since Keyg is closest to that value 201. | [6 empty] | Select orifice closest to value from Fig 3-16 for that line size.
The sample calculations, equations and spreadsheets presented herein were developed using examples published in the Engineering Data Book as published by the Gas Processor Suppliers Association as a service to the gas processing industry. All information and calculation formulae has been compiled and edited in cooperation with Gas Processors Association (GPA).
While every effort has been made to present accurate and reliable technical information and calculation spreadsheets based on the GPSA Engineering Data Book sample calculations, the use of such information is voluntary and the GPA and GPSA do not guarantee the accuracy, completeness, efficacy or timeliness of such information. Reference herein to any specific commercial product, calculation method, process, or service by trade-name, trademark, and service mark manufacturer or otherwise does not constitute or imply endorsement, recommendation or favoring by the GPA and/or GPSA.
The Calculation Spreadsheets are provided without warranty of any kind including warranties of accuracy or reasonableness of factual or scientific assumptions, studies or conclusions, or merchantability, fitness for a particular purpose or non-infringement of intellectual property.
In no event will the GPA or GPSA and their members be liable for any damages whatsoever (including without limitation, those resulting from lost profits, lost data or business interruption) arising from the use, inability to , reference to or reliance on the information in thes Publication, whether based on warranty, contract, tort or any other legal theory and whether or not advised of the possibility of such damages.
These calculation spreadsheets are provided to provide an “Operational level” of accuracy calculation based on rather broad assumptions (including but not limited to; temperatures, pressures, compositions, imperial curves, site conditions etc) and do not replace detailed and accurate Design Engineering taking into account actual process conditions, fluid properties, equipment condition or fowling and actual control set-point dead-band limitations.

## Example 3-2
Example 3-2 -- Calculate an approximate flow rate for the orifice using appropriate terms for eq. 3-4. | [6 empty] | Application 3-2 -- Calculate an approximate flow rate for the orifice using appropriate terms for eq. 3-4.
Given Data: | [6 empty] | Given Data:
Line Size, D | = | 6.605 | in |  |  |  | Line Size, D | = | 6.605 | in
Orifice Size, d | = | 3.5 | in |  |  |  | Orifice Size, d | = | 3.5 | in
Flange Taps | [6 empty] | Flange Taps
Flowing Temperature, Tf | = | 70 | °F |  |  |  | Flowing Temperature | = | 70 | °F
Flowing Pressure, Pf | = | 90 | psia |  |  |  | Flowing Pressure | = | 90 | psia
Differential, hw | = | 60 | in. of water |  |  |  | Differential | = | 60 | in. of water
Specific Gravity, G | = | 0.75 |  |  |  |  | Relative Density | = | 0.75
Keyg value from Fig 3-13 | = | 2646 |  |  |  |  | Keyg value from Fig 3-13 | = | 2646
From Fig. 3-3, | [6 empty] | From Fig. 3-3,
Qh | = | Keyg • Ftf • Fg • (hw • Pf)1/2 |  |  | from Equation 3-4 |  | Qh | = | Keyg • Ftf • Fg • (hw • Pf)1/2 |  |  | from Equation 3-4
Intermediate Calculations (not shown) | [6 empty] | Intermediate Calculations (not shown)
Ftf | = | [520 / (460 + Tf)]1/2 |  | = | 0.99052 | Fig 3-4 | Ftf | = | [520 / (460 + Tf)]1/2 |  | = | 0.99052 | Fig 3-4
Fg | = | [1.000/G]1/2 |  | = | 1.1547 | Fig 3-4 | Fg | = | [1.000/G]1/2 |  | = | 1.1547 | Fig 3-4
Qh | = | 2646 (0.9905) (1.1547) (hw • Pf)1/2 |  | = | 222392.3403 | cu ft/hr @ 14.73 psia and 60°F | Qh | = | 2646 (0.9905) (1.1547) (hw • Pf)1/2 |  | = | 222392.3403 | cu ft/hr @ 14.73 psia and 60°F
The sample calculations, equations and spreadsheets presented herein were developed using examples published in the Engineering Data Book as published by the Gas Processor Suppliers Association as a service to the gas processing industry. All information and calculation formulae has been compiled and edited in cooperation with Gas Processors Association (GPA).
While every effort has been made to present accurate and reliable technical information and calculation spreadsheets based on the GPSA Engineering Data Book sample calculations, the use of such information is voluntary and the GPA and GPSA do not guarantee the accuracy, completeness, efficacy or timeliness of such information. Reference herein to any specific commercial product, calculation method, process, or service by trade-name, trademark, and service mark manufacturer or otherwise does not constitute or imply endorsement, recommendation or favoring by the GPA and/or GPSA.
The Calculation Spreadsheets are provided without warranty of any kind including warranties of accuracy or reasonableness of factual or scientific assumptions, studies or conclusions, or merchantability, fitness for a particular purpose or non-infringement of intellectual property.
In no event will the GPA or GPSA and their members be liable for any damages whatsoever (including without limitation, those resulting from lost profits, lost data or business interruption) arising from the use, inability to , reference to or reliance on the information in thes Publication, whether based on warranty, contract, tort or any other legal theory and whether or not advised of the possibility of such damages.
These calculation spreadsheets are provided to provide an “Operational level” of accuracy calculation based on rather broad assumptions (including but not limited to; temperatures, pressures, compositions, imperial curves, site conditions etc) and do not replace detailed and accurate Design Engineering taking into account actual process conditions, fluid properties, equipment condition or fowling and actual control set-point dead-band limitations.

## Example 3-3
Example 3-3 -- Calculate the daily gas flow through a 1-inch orifice in a nominal 3-inch pipe. The gas relative density is 0.60, the flowing temperature is 60°F, and the pressure upstream of the orifice is 5 inches Hg. The published ID of a 3-inch pipe is 3.068 in. | [8 empty] | Application 3-3 -- Calculate the daily gas flow through a 1-inch orifice in a nominal 3-inch pipe. The gas relative density is 0.60, the flowing temperature is 60°F, and the pressure upstream of the orifice is 5 inches Hg. The published ID of a 3-inch pipe is 3.068 in.
Given Data: | [8 empty] | Given Data:
Relative Density | = | 0.6 | [6 empty] | Relative Density | = | 0.6
Orifice Size, d | = | 1 | in | [5 empty] | Orifice Size, d | = | 1 | in
Line Size, ID | = | 3.068 | in | [5 empty] | Line Size, ID | = | 3.068 | in
Flowing Temperature | = | 60 | °F | [5 empty] | Flowing Temperature | = | 60 | °F
Flowing Pressure | = | 5 | in Hg | [5 empty] | Flowing Pressure | = | 5 | in Hg
Q | = | 16,330 (1 + β4) (d)2 ((H(29.32 + 0.3H))1/2 • Ftf • Cg |  |  |  |  | Eq. 3-14 |  | Q | = | 16,330 (1 + β4) (d)2 ((H(29.32 + 0.3H))1/2 • Ftf • Cg |  |  |  |  | Eq. 3-14
Intermediate Calculations (not shown) | [8 empty] | Intermediate Calculations (not shown)
Ftf | = | [520 / (460 + Tf)]1/2 |  | = | 1 | Eq. 3-15 |  |  | Ftf | = | [520 / (460 + Tf)]1/2 |  | = | 1 | Eq. 3-15
Cg | = | [0.60/G]1/2 |  | = | 0.92582 | Eq. 3-16 |  |  | Cg | = | [0.60/G]1/2 |  | = | 0.92582 | Eq. 3-16
Q | = | 16,330 (1 + (1/3.068)4) (1.0)2 ((5(29.32 + 0.3)(5))1/2 = |  |  |  |  | 189796.6464 | std cu ft per day | Q | = | 16,330 (1 + (1/3.068)4) (1.0)2 ((5(29.32 + 0.3)(5))1/2 = |  |  |  |  | 189796.6464 | std cu ft per day
The sample calculations, equations and spreadsheets presented herein were developed using examples published in the Engineering Data Book as published by the Gas Processor Suppliers Association as a service to the gas processing industry. All information and calculation formulae has been compiled and edited in cooperation with Gas Processors Association (GPA).
While every effort has been made to present accurate and reliable technical information and calculation spreadsheets based on the GPSA Engineering Data Book sample calculations, the use of such information is voluntary and the GPA and GPSA do not guarantee the accuracy, completeness, efficacy or timeliness of such information. Reference herein to any specific commercial product, calculation method, process, or service by trade-name, trademark, and service mark manufacturer or otherwise does not constitute or imply endorsement, recommendation or favoring by the GPA and/or GPSA.
The Calculation Spreadsheets are provided without warranty of any kind including warranties of accuracy or reasonableness of factual or scientific assumptions, studies or conclusions, or merchantability, fitness for a particular purpose or non-infringement of intellectual property.
In no event will the GPA or GPSA and their members be liable for any damages whatsoever (including without limitation, those resulting from lost profits, lost data or business interruption) arising from the use, inability to , reference to or reliance on the information in thes Publication, whether based on warranty, contract, tort or any other legal theory and whether or not advised of the possibility of such damages.
These calculation spreadsheets are provided to provide an “Operational level” of accuracy calculation based on rather broad assumptions (including but not limited to; temperatures, pressures, compositions, imperial curves, site conditions etc) and do not replace detailed and accurate Design Engineering taking into account actual process conditions, fluid properties, equipment condition or fowling and actual control set-point dead-band limitations.

## Example 3-4
Example 3-4 -- Calculate an approximate orifice size for the given flow rate and line size | [6 empty] | Application 3-4 -- Calculate an approximate orifice size for the given flow rate and line size
Given Data: | [6 empty] | Given Data:
Line Size, D | = | 3.068 | in |  |  |  | Line Size, D | = | 3.068 | in
Flange Taps | [6 empty] | Flange Taps
Relative Density @ 60 °F | = | 0.69 |  |  |  |  | Relative Density @ 60 °F | = | 0.69
Flowing Temperature | = | 40 | °F |  |  |  | Flowing Temperature | = | 40 | °F
Flow Rate | = | 3400 | gal. per hr |  |  |  | Flow Rate | = | 3400 | gal. per hr
|  | @ 14.73 | psia, 60 ºF | [5 empty] | @ 14.73 | psia, 60 ºF
Desired Differential | = | 50 | in. of water |  |  |  | Desired Differential | = | 50 | in. of water
To determine the approximate orifice size required, the corresponding Keyl is calculated using appropriate terms of Equation 3-7. | [6 empty] | To determine the approximate orifice size required, the corresponding Keyl is calculated using appropriate terms of Equation 3-7.
Qh | = | Keyl • Fgt • (hw)1/2 |  |  | From Equation 3-7 |  | Qh | = | Keyl • Fgt • (hw)1/2 |  |  | From Equation 3-7
or Keyl | = | Qh / (Fgt • (hw)1/2 |  |  |  |  | or Keyl | = | Qh / (Fgt • (hw)1/2
Fgt | = | [1.0057/(G)1/2] • [Gf /Gl]1/2 |  | = | 1.2193 | Fig 3-4 | Fgt | = | [1.0057/(G)1/2] • [Gf /Gl]1/2 |  | = | 1.2193 | Fig 3-4
Keyl | = | 3400 |  | = | 394.3536 |  | Keyl | = | 3400 |  | = | 394.3536
|  | 1.2193 • (50)1/2 | [6 empty] | 1.2193 • (50)1/2
The sample calculations, equations and spreadsheets presented herein were developed using examples published in the Engineering Data Book as published by the Gas Processor Suppliers Association as a service to the gas processing industry. All information and calculation formulae has been compiled and edited in cooperation with Gas Processors Association (GPA).
While every effort has been made to present accurate and reliable technical information and calculation spreadsheets based on the GPSA Engineering Data Book sample calculations, the use of such information is voluntary and the GPA and GPSA do not guarantee the accuracy, completeness, efficacy or timeliness of such information. Reference herein to any specific commercial product, calculation method, process, or service by trade-name, trademark, and service mark manufacturer or otherwise does not constitute or imply endorsement, recommendation or favoring by the GPA and/or GPSA.
The Calculation Spreadsheets are provided without warranty of any kind including warranties of accuracy or reasonableness of factual or scientific assumptions, studies or conclusions, or merchantability, fitness for a particular purpose or non-infringement of intellectual property.
In no event will the GPA or GPSA and their members be liable for any damages whatsoever (including without limitation, those resulting from lost profits, lost data or business interruption) arising from the use, inability to , reference to or reliance on the information in thes Publication, whether based on warranty, contract, tort or any other legal theory and whether or not advised of the possibility of such damages.
These calculation spreadsheets are provided to provide an “Operational level” of accuracy calculation based on rather broad assumptions (including but not limited to; temperatures, pressures, compositions, imperial curves, site conditions etc) and do not replace detailed and accurate Design Engineering taking into account actual process conditions, fluid properties, equipment condition or fowling and actual control set-point dead-band limitations.

## Example 3-5
Example 3-5 -- Calculate a liquid flowrate for the given orifice setting. | [6 empty] | Application 3-5 -- Calculate a liquid flowrate for the given orifice setting.
Line Size, D | = | 8.071 | in |  |  |  | Line Size, D | = | 8.071 | in
Orifice Size, d | = | 4 | in |  |  |  | Orifice Size, d | = | 4 | in
Flange Taps | [6 empty] | Flange Taps
Flowing Temperature | = | 80 | °F |  |  |  | Flowing Temperature | = | 80 | °F
Flowing Pressure | = | 90 | psia |  |  |  | Flowing Pressure | = | 90 | psia
Differential | = | 36 | in. of water |  |  |  | Differential | = | 36 | in. of water
Relative Density | = | 0.63 |  |  |  |  | Relative Density | = | 0.63
To determine the flow rate from Fig. 3-4, | [6 empty] | To determine the flow rate from Fig. 3-4,
Qh | = | Keyl • Fgt • (hw)1/2 |  |  | From Equation 3-7 |  | Qh | = | Keyl • Fgt • (hw)1/2 |  |  | From Equation 3-7
Fgt | = | [1.0057/(G)]1/2 • [Gf /Gl]1/2 |  | = | 1.256 | Fig 3-4 | Fgt | = | [1.0057/(G)]1/2 • [Gf /Gl]1/2 |  | = | 1.256 | Fig 3-4
The value of Keyl from Fig. 3-13 is 3345 for an 8.071 in. line | [6 empty] | The value of Keyl from Fig. 3-13 is 3345 for an 8.071 in. line
with a 4.0 in. orifice. The value of Fgt is calculated from the Fig. 3-4 equation | [6 empty] | with a 4.0 in. orifice. The value of Fgt is calculated from the Fig. 3-4 equation
Therefore, | [6 empty] | Therefore,
Qh | = | 3345 • 1.2560 • (36)1/2 |  | = | 25206.9673 | gal/hr | Qh | = | 3345 • 1.2560 • (36)1/2 |  | = | 25206.9673 | gal/hr
The sample calculations, equations and spreadsheets presented herein were developed using examples published in the Engineering Data Book as published by the Gas Processor Suppliers Association as a service to the gas processing industry. All information and calculation formulae has been compiled and edited in cooperation with Gas Processors Association (GPA).
While every effort has been made to present accurate and reliable technical information and calculation spreadsheets based on the GPSA Engineering Data Book sample calculations, the use of such information is voluntary and the GPA and GPSA do not guarantee the accuracy, completeness, efficacy or timeliness of such information. Reference herein to any specific commercial product, calculation method, process, or service by trade-name, trademark, and service mark manufacturer or otherwise does not constitute or imply endorsement, recommendation or favoring by the GPA and/or GPSA.
The Calculation Spreadsheets are provided without warranty of any kind including warranties of accuracy or reasonableness of factual or scientific assumptions, studies or conclusions, or merchantability, fitness for a particular purpose or non-infringement of intellectual property.
In no event will the GPA or GPSA and their members be liable for any damages whatsoever (including without limitation, those resulting from lost profits, lost data or business interruption) arising from the use, inability to , reference to or reliance on the information in thes Publication, whether based on warranty, contract, tort or any other legal theory and whether or not advised of the possibility of such damages.
These calculation spreadsheets are provided to provide an “Operational level” of accuracy calculation based on rather broad assumptions (including but not limited to; temperatures, pressures, compositions, imperial curves, site conditions etc) and do not replace detailed and accurate Design Engineering taking into account actual process conditions, fluid properties, equipment condition or fowling and actual control set-point dead-band limitations.

## Example 3-6
Example 3-6 -- The following information was taken from a meter proving report and is used to calculate a meter factor with which to correct the routine meter recorded volumes. | [5 empty] | Application 3-6 -- The following information was taken from a meter proving report and is used to calculate a meter factor with which to correct the routine meter recorded volumes.
Given Data: | [5 empty] | Given Data:
Product | = | Propane-Butane Mix |  |  |  | Product | = | Propane-Butane Mix
Relative Density | = | 0.544 |  |  |  | Relative Density | = | 0.544
Totalizer Pulses Per Barrel, K | = | 13188 |  |  |  | Totalizer Pulses Per Barrel, K | = | 13188
Base Prover Volume | = | 2.1612 | bbl |  |  | Base Prover Volume | = | 2.1612 | bbl
Vapor Pressure | = | 115 | psig |  |  | Vapor Pressure | = | 115 | psig
Prover Size | = | 12 | in. pipe with |  |  | Prover Size | = | 12 | in. pipe with
|  | 0.375 | in. wall |  |  |  |  | 0.375 | in. wall
Vertical Double Case Meter | [5 empty] | Vertical Double Case Meter
The volume of a steel prover must first be determined then the characteristics of the fluid are used to correct the fluid to standard conditions using the corrected prover volume | [5 empty] | The volume of a steel prover must first be determined then the characteristics of the fluid are used to correct the fluid to standard conditions using the corrected prover volume
Prover Corrections: | [5 empty] | Prover Corrections:
1. Base prover volume |  |  | = | 2.1612 | bbl | 1. Base prover volume |  |  | = | 2.1612 | bbl
2. The correction for steel expansion at at 77.0°F, Cts (Fig. 3-23) |  |  | = | 1.0003 |  | 2. The correction for steel expansion at at 77.0°F, Cts (Fig. 3-23) |  |  | = | 1.0003
3. The correction for the pressure effect on the steel prover, Cps (Fig. 3-24) |  |  | = | 1.0004 |  | 3. The correction for the pressure effect on the steel prover, Cps (Fig. 3-24) |  |  | = | 1.0004
4. The correction for the effect of temperature on the prover fluid, Ctl (Reference No. 4) |  |  | = | 0.978 |  | 4. The correction for the effect of temperature on the prover fluid, Ctl (Reference No. 4) |  |  | = | 0.978
5. The correction for the effect of pressure on the prover fluid, Cpl (Reference No. 2) |  |  | = | 1.0078 |  | 5. The correction for the effect of pressure on the prover fluid, Cpl (Reference No. 2) |  |  | = | 1.0078
6. Combined prover corrections | = | Cts • Cps • Ctl • Cpl | = | 0.98632 |  | 6. Combined prover corrections | = | Cts • Cps • Ctl • Cpl | = | 0.98632
7. Gross standard volume of the prover | = | .9863 x 2.1612 bbl | = | 2.1316 | bbl | 7. Gross standard volume of the prover | = | .9863 x 2.1612 bbl | = | 2.1316 | bbl
Metered fluid corrections: | [5 empty] | Metered fluid corrections:
1. IV | = | 28631/13188 | = | 2.171 | bbl | 1. IV | = | 28631/13188 | = | 2.171 | bbl
2. The correction for the effect of temperature on the metered volume, Ctl (Reference No. 4) |  |  | = | 0.9789 |  | 2. The correction for the effect of temperature on the metered volume, Ctl (Reference No. 4) |  |  | = | 0.9789
3. The correction for the effect of pressure on the metered volume, Cpl (Reference No. 2) |  |  | = | 1.008 |  | 3. The correction for the effect of pressure on the metered volume, Cpl (Reference No. 2) |  |  | = | 1.008
4. Indicated standard volume of meter | = | Ctl • Cpl • 2.1710 | = | 2.1422 | bbl | 4. Indicated standard volume of meter | = | Ctl • Cpl • IV | = | 2.1422 | bbl
5. MF = gross standard volume of prover/indicated std volume of meter | = | 2.1316 | = | 0.99507 |  | 5. MF = gross standard volume of prover/indicated std volume of meter | = | 2.1316 | = | 0.99507
|  | 2.1422 | [5 empty] | 2.1422
The sample calculations, equations and spreadsheets presented herein were developed using examples published in the Engineering Data Book as published by the Gas Processor Suppliers Association as a service to the gas processing industry. All information and calculation formulae has been compiled and edited in cooperation with Gas Processors Association (GPA).
While every effort has been made to present accurate and reliable technical information and calculation spreadsheets based on the GPSA Engineering Data Book sample calculations, the use of such information is voluntary and the GPA and GPSA do not guarantee the accuracy, completeness, efficacy or timeliness of such information. Reference herein to any specific commercial product, calculation method, process, or service by trade-name, trademark, and service mark manufacturer or otherwise does not constitute or imply endorsement, recommendation or favoring by the GPA and/or GPSA.
The Calculation Spreadsheets are provided without warranty of any kind including warranties of accuracy or reasonableness of factual or scientific assumptions, studies or conclusions, or merchantability, fitness for a particular purpose or non-infringement of intellectual property.
In no event will the GPA or GPSA and their members be liable for any damages whatsoever (including without limitation, those resulting from lost profits, lost data or business interruption) arising from the use, inability to , reference to or reliance on the information in thes Publication, whether based on warranty, contract, tort or any other legal theory and whether or not advised of the possibility of such damages.
These calculation spreadsheets are provided to provide an “Operational level” of accuracy calculation based on rather broad assumptions (including but not limited to; temperatures, pressures, compositions, imperial curves, site conditions etc) and do not replace detailed and accurate Design Engineering taking into account actual process conditions, fluid properties, equipment condition or fowling and actual control set-point dead-band limitations.